Estimation of the Instrument Profile for LASCO C1 Images 
Paul Reiser
Email: reiser@susim.nrl.navy.mil
Phone: (202) 404 1413

Definition of the Instrument Profile 

The Fabry-Perot (FP) interferometer and its associated blocking filters transmits light in a narrow band of wavelengths. The normalized transmission as a function of wavelength is called the instrument profile of the instrument. The instrument profile of an ideal FP is given by a Lorentzian Profile: 

        L(w) = g/(pi(w^2+g^2))                Eq(1)

where w is wavelength, L(w) is the instrument profile and the full width at half maximum (FWHM) is given by: 

        fl == 2g                              Eq(2)

where fl is the FWHM. (Note that the integral of L(w) over wavelength is normalized to unity.) 

In the case of the LASCO C1 instrument, it has been found that the instrument profile is best represented by a Voigt function which is a convolution of the Lorentzian profile and a Gaussian profile. The Gaussian profile may be written as: 

        G(w) = exp(-w^2/(2 s^2))              Eq(3)

               ----------------

                  s sqrt(2 pi)

where G(w) is the profile and s is the variance of the Gaussian. Note that, as in the case of the Lorentzian, the Gaussian profile is also normalized to unity. Note also that the FWHM of the Gaussian profile is given by fg where: 

        fg = 2s sqrt(2 ln(2))                 Eq(4)

Measurement of the Instrument Profile 

Measurements of the Lorentz and Gaussian profile parameters for the LASCO C1 instrument have been made and the table below lists the values (in Angstroms) for the five different wavelength regimes represented by the five individual blocking filters: 

              Blocker          fl (A)          fg (A)           f (A)

        5303 FeXIV Green Line   0.064           0.537           0.572

        5694 CaXV  Blue Line    0.242           0.424           0.571

        5890 NaI   Yellow Line  0.161           0.562           0.654

        6374 FeX   Red Line     0.392           0.535           0.781

        6563 Hydrogen Alpha     0.089           0.947           0.995

                                   Table 1

Where fl is the Lorentzian FWHM, fg is the Gaussian FWHM, and f is the Voigt FWHM. These values were arrived at by minimising the RMS error in the difference between the measured profile and the NSO[1] solar spectrum convolved with the Voigt instrument function. 


References
[1] Kurucz, R.L., Furnlid, I., Brault, J., Testerman, L, ``Solar Flux Atlas from 296 to 1300 nm'', National Solar Observatory Atlas No. 1, June 1984.

Paul Reiser
Sep 26, 1998
Apr 08, 1999 

